## Slide 1: Main Pod Code Guide
V4
March 2025

## Slide 2: For Existing Users:
Re-download main_pod_code_2025.mat
Re-download Addtl. Matlab Functions folder
Re-download Logs folder
Delete any existing .mat pod files in colocation or field folders
  Code will re-convert them to .mat reflecting new formats

## Slide 3: Calibration Methods

## Slide 4: What is Individual Calibration?
Each pod has its own distinct calibration mode
E.g. for these 6 pods, 6 unique sets of calibration coefficients

## Slide 5: What is universal (multi-pod) calibration? One-Cal
Only one pod needs to be co-located with a reference instrument
One calibration model is generated, then applied to all 6 pod
  Only 1 calibration model total

## Slide 6: What is universal (multi-pod) calibration? One-Hop
  One pod is normalized, then treated as a “reference” instrument in a secondary colocation to normalize the rest of the pods
    6 calibration models total:
      1 derived from reference instrument
      5 from using the calibrated pod as a “reference” instrument

## Slide 7: Getting Set up in Matlab

## Slide 8: Matlab Version
Optimized for use with R2021b
Can still use with newer versions, but may experience bugs

## Slide 9: Additional Matlab toolboxes needed
The following MatLab toolboxes are required to run the main script and associated plotting functions:
  Deep Learning Toolbox
  Statistics and Machine Learning Toolbox
  Curve Fitting Toolbox
  Optimization Toolbox
  Econometrics Toolbox

## Slide 10: Additional Matlab toolboxes needed
Gramm is required for most plotting functions
  Download here: https://www.mathworks.com/matlabcentral/fileexchange/54465-gramm-data-visualization-toolbox
  Download as a zip file
  Unzip on your computer
  In the MATLAB folder on your computer (auto-generated when you download Matlab), add the contents of the gramm folder as shown here:

## Slide 11: Folders You’ll Need to Run the Code

## Slide 12: Folders You’ll Need to Run the Code
Need certain folders saved to your computer that the code will know to pull files from:
Addtl Matlab Functions
  Version must be October 2021 or later
Colocation
Colocation2
  If using one-hop, otherwise don’t need
Field
Logs
These go in the same folder as the code you want to run

## Slide 13: Folders You’ll Need to Run the Code > Addtl Matlab Functions
  Houses all of the sub-functions the code uses
  Don’t edit this unless you want to create a new calibration model, make a new type of plot, etc.

## Slide 14: A note on Addtl Matlab Functions
  Most sub-folders have an example function to use as a template when writing your own
  Titled: a_generic(FunctionType)

## Slide 15: Folders You’ll Need to Run the Code > Colocation
Split into 2 sub-folders:
Pod:
  Place raw .txt files here for your co-location data
  Code will automatically turn any new .txt files to .mat files
  Need at least 2 pods in here for the code to run
    If universal cal (only one pod), put a dummy file in
Reference:
  Place the reference data here in .csv format (Excel)
  Needs 2 columns: datetime & [pollutant name]
  Code will automatically convert it to .mat file
  If coarse & fine PM, need to manually split into PM10, 5, etc.

## Slide 16: Folders You’ll Need to Run the Code > Colocation2
Only for one-hop
You only need to create the “Pods” folder
Pods:
  Place raw .txt files here for your secondary co-location data
  Code will automatically turn any new .txt files to .mat files
  Need at least 2 pods in here for the code to run
Reference:
  Will be automatically generated by the code
Before running code
After running code

## Slide 17: Folders You’ll Need to Run the Code > Field
  Place raw .txt files here for your field / deployment data
  Code will automatically turn any new .txt files to .mat files
    If changes are made to raw .txt files, delete the corresponding .mat files so code regenerates those

## Slide 18: A note on pod .txt files
  If more than one file for each pod during colocation or field, best to combine them
    Combine multiple colocation files or multiple field files
    Don’t combine colocation with field
  Can do this using windows command line
    Add multiple field files for the same pod to one folder
    Open command prompt by typing “cmd” into windows search bar
    Type “cd” then the directory of your new folder
    Copy & paste: copy *.txt newfile.txt
    This will create an additional file with all combined

## Slide 19: Folders You’ll Need to Run the Code > Logs
  Contains 2 .csv files (Excel)
  Deployment Log
  Pod Inventory

## Slide 20: Logs > Deployment Log.csv
Needs to have the first 5 headings
  The rest are good to have for the database but not required for code to run
Separate entries for each pod’s co-location & deployment
PodName: needs to be UPODB2, YPODB2 (not just B2)
TimeZone: difference from UTC (-7 or -8 for CO)
ReferenceFile:
  For co-location, the name of the file in Colocation > Reference folder
  For 1-hop: Colocation2
  For deployment/field data: just put NA

## Slide 21: Logs > Pod Inventory.csv
Use these exact column headings
Timezone: leave as 0
  Assumes UTC as timezone
Need to have an entry for each pod used
  Must be UPODB2, YPODB2 (not just B2)
Update this from previous versions – some names have changed

## Slide 22: Logs > Pod Inventory.csv
VarNames: lists the names you want for each column in the pod text file
  E.g. you’ll add “ptPM1, ptPM2.5, ptPM10” as variables 9, 10, 11 etc.
VarTypes: lists the type of input this column will be
  E.g you’ll add “double, double, double” as variables 9, 10, 11 etc.

## Slide 23: SettingsSet

## Slide 24: settingsSet
Modify settings here (begins line 23)
All variables to be edited start with settingsSet
The rest of the code should remain untouched
You don’t edit anything below here
(end of settingsSet)
Starts line 29

## Slide 25: settingsSet > calibration mode (calMode)
Determines if individual or either type of universal cal is run
  0 – individual calibration
  1 – universal calibration using one-cal method
  2 – universal calibration using one-hop method

## Slide 26: settingsSet > universal Pod (uniPod)
If using either universal calibration method: which pod you want to use to calibrate the rest
Can leave blank if doing individual calibration

## Slide 27: settingsSet > portions of code to run
For field application, applyCal must be true
If you want text files to save to the database, database must be true

## Slide 28: settingsSet > datetime
Default settings are what works best for Y-Pod data & later
Only change if using U-Pod data or older

## Slide 29: settingsSet > Averaging
Chose time averaging (minutely, hourly, etc.) based on resolution of reference data
Can also choose “smoothing” method for making data minutely, hourly, etc.

## Slide 30: settingsSet > data extraction
Choose reference gas
  Must match the column name in the reference folder
  Do one reference at a time! E.g. if you want to look at both CH4 & CO, run code twice instead of listing both in one run
Choose which pod pollutant sensors to import
Choose which pod environmental sensors to import
Only import columns you need to quantify pollutant of interest
  Importing them all slows things down
  Can confuse the calibration model

## Slide 31: settingsSet > function names
Choose what pre-processing to use for reference data
  In this section, choose whether to use voltage, resistance, or normalized resistance
Choose what pre-processing to use for pod data
  From Addtl. Matlab Functions > preprocessfunctions for both
Choose which calibration model(s) to run
  From Addtl. Matlab Functions > Calibration Models
Function names must appear in “Addtl. Matlab Functions” folder

## Slide 32: settingsSet > function names
If using field application (apply cal), choose one model to use on field data
  From Addtl. Matlab Functions > Calibration Models
This same model needs to be used in modelList
Can use as many models as you want in modelList, but only one in fieldModel

## Slide 33: settingsSet > function names
Next few functions deal with splitting co-location data into calibration & validation
valList – what method to split into cal & val
  Needs to be a filename from Addtl. Matlab Functions > ValidationFunctions
  Select data randomly, group based on time of day, temperature, etc.
nFolds & nFolds rep – what percentage of data is used for validation
  5 folds = 20%, 3 folds = 33%, etc.

## Slide 34: settingsSet > function names
statsList – which statistics to calculate for co-location fits
  Addtl. Matlab Functions > StatFunctions
  Need the 3 listed here to run ‘basicStatsBoxplots’
plotsList – plot stats, timeseries, etc. for co-location fits
  Addtl. Matlab Functions > PlottingFunctions

## Slide 35: settingsSet > field
DO NOT EDIT THIS ENTRY!! ALWAYS LEAVE SET AS ‘no’
The code will automatically run any pod pre-processing to both the colocation and field data. This is for any pre-processing that should be applied to the field data (rare case).

## Slide 36: settingsSet > ispc
The last entry in the settingsSet
If you’re on a pc, leave as is
If on a mac, change to ‘no’

## Slide 37: Calibration Models

## Slide 38: Recommended CalibrationModels
For O3 & CO2: ‘rf_plain’ or ‘rf_plainTe’
For CH4 & HCHO: ‘rf_HCHO’ or ‘rf_HCHO_Te’
For NO2: have not used recently, but recommend using both the NO2 & O3 sensors for better fits. Can also try including additional electrochemical sensor (CO).
Rf are all random forest models
For shorter deployments, the elapsed time term (Te) may cause models to overfit, which leads to issues with field data – try both & look carefully at the field outputs!

## Slide 39: A note on joannaNN and ANNs
joannaNN is a base ANN model
  Can tweak parameters (some examples shown below) in this function
  Or create other versions named accordingly
    e.g. joannaNN_10n_lm_100e, joannaNN_ratioSens1

## Slide 40: A note on joannaNN and ANNs
joannaNN just uses whichever sensor columns you selected
  From podSensors & envSensors
    E.g. Fig2600, Fig2602, temperature, humidity
Add interaction terms by calling the corresponding cal model instead
  E.g. products & ratios of metal oxide sensors, temp & hum interactions

## Slide 41: Recommended Plots
*these are all for colocation data. Field plots are generated outside of this code!

## Slide 42: PlottingFunctions > acfPlot
Autocorrelation function
More info here: https://www.geeksforgeeks.org/autocorrelation/
For our purposes, helps us figure out if our time zones between the pods and reference instrument are misaligned
If lag = 0 or lag < ~30 min, time zone is ok
If lag > 30 min, try adjusting time zone in deployment log
2 examples of reliable data (small or no lags)

## Slide 43: PlottingFunctions > basicStatsBoxPlots
Plots the statistics for goodness of fit for each pod
  Since each pod is split into folds for testing & training, will have a range of values for each pod
  We usually report out the median of these values (center line on each box)
Default stats included: R2, RMSE (root mean square error) MBE (mean bias error)
  More options of stats available in Addtl Matlab Functions > StatFunctions

## Slide 44: PlottingFunctions > timeseriesPlot
Shows timeseries of the reference data on top and the calibrated sensor data on bottom for comparison purposes

## Slide 45: PlottingFunctions > vsReferencePlot
Plots the reference data on the X axis and our calibrated pod data on the Y axis
If our fit was perfect, it would be a perfectly straight line with equation y = x (y = mx + b, where slope m =1 and intercept b = 0)
Deviation from the 1-1 line helps us determine where the model over or under estimates
In this example, the trailing points let us know that the model overestimates for small CO values and underestimates for high CO values

## Slide 46: PlottingFunctions > XYCorrelations
Plots each included pod variable against each other to help us look for trends or issues
The straight lines are where the variable is being compared against itself

## Slide 47: Body of Code

## Slide 48: List of Functions
Jake had compiled a list of many functions available in Addtl. Matlab Functions with a description of each
There are others I added not included here – ask!

## Slide 49: Begin Body of Code
This is where everything happens!
Shouldn’t need to edit anything below here

## Slide 50: Outputs

## Slide 51: Outputs > General
Naming convention: Pod ID_output description_model used
  Ex. YPODB2_CalValEstimates_ratioSens1.mat
run_settings.mat
  View which settings you selected in settingsSet
Any selected plots saved as .fig (Matlab Figure)
  There will be an .m file with the same name as each figure
  Contains the data needed to open the file

## Slide 52: Outputs > Calibration
CalVal_Estimates: structure containing data from cal/val folds
  Can extract if needed but usually not necessary
Colocation_Estimate: model applied to full set of colocation data
  Very useful for additional plotting & analysis
Fit_t: timestamps used for colocation data
Fit_X: preprocessed pod data from colocation
Fit_Y: reference data aligned to match with pod timestamps
fittedMdl: actual calibration model used

## Slide 53: Outputs > Field
Field_Estimate: model applied to full set of field data
Field_t – timestamps used for field data
Field_X – preprocessed pod data from field

## Slide 54: Outputs > Reference
If using one-hop, a “Reference” folder will appear in the “Calibration2” folder
  A new one is generated on each run with the datetime string it was created
Y_ref: “reference” estimates for the universal pod during the secondary colocation
yt: matching timestamps for the secondary colocation

## Slide 55: Outputs > Database
If settingsSet.database is true, generates .txt files to be saved to database
  One of each for each pod
Colocation: pod data, estimate, timestamps, and reference data
Colocation2: same info but for one-hop & calibrating pod only
Field: pod data, estimate, and timestamps for field data
settingsSet: all run settings as a text file
deployLog: entire deployment log in text file format

## Slide 56: Key Changes from Previous Versions

## Slide 57: Key Changes from Version 3.3
“regtype” no longer included in settingsSet
  Previously, used had to note the type of regression used in code – either linear regression, ANN, RF, etc.
  New handling relies on calibration model naming convention being followed
    If calibration model name includes ‘NN’ or ‘RF’, will handle as appropriate

## Slide 58: Key Changes from (older) Previous Versions
Calibration & field all in one code!
Compatible with different types of machine learning (linreg, ann, universal)
One version of plotting functions, stat functions, etc. used
  Names & structures have been edited
  Update Addtl. Matlab Functions to have the correct version
Structure of calibration models has changed
  Older file versions of anything in Addtl. Matlab Functions > Calibration Models will be incompatible
  See next slide for more info

## Slide 59: Added functionality by adding additional cases to EACH calibration model
Originally 3 cases:
Generate calibration
Apply calibration to cal/val folds
Report out on calibration
Now requires 5 cases:
Generate calibration
Apply calibration to cal/val folds
Apply calibration to ALL colocation data
Apply calibration to field data
Report out on calibration
Make sure you’ve updated “Addtl Matlab Functions” folder to have the correct versions of the cal models!!

## Slide 60: Note on Calibration Models
Only regression models have been updated for this version
  We only use regression for determining concentration estimates
  Regression has a continuous output
    E.g. estimate methane concentration, which can be any number (infinite possible outcomes)
Classification models still available in archive
  Would need the format updated as shown on previous slide in order to use
  Classification has a discrete output
    E.g determine whether a measurement represents oil & gas activity or not (2 possible outcomes)